1. Contents

Section Topic                                                      Page

01.    Table of Contents                                         Page 1

02.    Introduction                                            Pages 2-3

03.    Appearance and Key Functions of the meter    				Page	   4

04     Alternate Site Testing  								     Pages 5-8

05     Safety Information                               	       Page 9

06.    Quick Testing Instructions                               Pages 10-11

Preparing for Testing

07.    Installing the Battery    									Page	  12

08.	  Setting the Audio Voice and Language 	                   Pages 13-14

09.    Settings (Date/Time/Alarm/Measuring Unit and

Memory Deletion ) and LCD Screen Introduction            Pages 15-19

10.    Running a Control Solution Test                          Pages 20-21

Blood Testing

11.    Running a Blood Glucose Test							Pages	22-24

12.    Recalling the Memory and Viewing the Average   		Pages	25-28

13.    Expected Values Reference                 				Page 28

14.    Transmitting Results          							Page	   29

Maintenance

15.    Maintenance of Your System								Page	   30

16.    Troubleshooting                                        Page 31-34

17.    Customer Service                                        Page 34

18.    Limitations                                             Page 35

Important Additional Information

19.    Specifications											Page	  36

20.    Ordering Supplies                                        Page 37

21.    Warranty Information                                     Page 38

NOTE: For information about the name of the manufacturer of the lancing device

and the lancets, please refer to the package inserts that came with your starter kits.

2. Introduction

Intended Use/ Indications for Use

The Solo-V2TM Blood Glucose Monitoring System is intended for use in the quantitative measurement of glucose in fresh capillary whole blood from the finger and the forearm. It is intended for use by healthcare professionals and people with diabetes mellitus at home and as an aid in monitoring the effectiveness of a diabetes control program. The Solo- V2TM Blood Glucose Monitoring System is not intended for the diagnosis of or screening for diabetes mellitus, nor for use with neonates.

The alternative site testing (AST) in this system can only be used during steady-state blood glucose conditions.

This system contains a speaking function that provides an audible message of test results for users with low vision.

Contents of the System

This User Guide has been prepared to describe the proper use of Solo-V2TM Blood Glucose Monitoring System. Please read this User Guide and the Package Insert that is accompanied with the Solo-V2TM Glucose Test Strips before testing. The system is available as a kit. You can purchase the system supplies from your system provider. Please tell your system provider the 7 digit system reference number marked on the packaging box when you purchase supplies.

NDC#: 8611-501001

The starter kit of Solo-V2TM Blood Glucose Monitoring System includes the following items.

1. Solo-V2TM  Glucose Meter with (2) AAA batteries

2. User Guide

3. Log Book

4. Carrying Case

5. Solo-V2TM Glucose Test Strips

6. Solo-V2TM Control Solution (Low)

7. Solo V2 Sterile Lancets

8. Solo V2 Lancing Device

9. Package Inserts for Test Strips, Control Solution, and Lancets/ Lancing Device

A single level control solution (Low) is provided in the “Starter Kit”, and you may purchase High or Low control solution that you need from your system provider.

The Solo-V2TM Glucose Meter uses Solo-V2TM Test Strips only. Neither the meter nor the test strips will work when used with any other brand of glucose products.

3. Appearance and Key functions of the Meter

1. Test strip insertion port – Located at the top of the meter and easily identified by a notch. Meter will automatically turn on when a strip is inserted.

2. LCD Display - Guides the user through the test using symbols and simple messages.

3. Left Function Key – Located just below the LCD screen on the front of the meter are two function buttons. The Left Function key will power ON/OFF, also for memory recalling mode.

4. Right Function Key – Located just below the LCD screen and directly to the right of the Left Function key. This key confirms settings, allows ability to view averages in memory mode, and flags particular test results.

5. Speaker Slot – Located directly below the Left and Right Function keys on the front of the meter and can be identified by the 3 vertical slots in the meter. Place where the audible output is produced.

6. Up Button – Located towards the top of the right side of the meter. It is the first button on the right side of the meter starting from the top of the meter going downward. Increase the value of the current setting and increases the voice volume.

7. Down Button – Located on the right side of the meter directly below the Up button. Decreases the value of the current setting and decreases the voice volume.

8. Language Selection Button – Located directly below the Down button on the right side of the meter. This button does not protrude from the meter as much as the Up and Down buttons and is a round button. It allows user to change between audible languages.

9. Mini USB Port – Located directly below the Language Selection button on the right side of the meter. This port is where you connect cable between meter and PC for data transmitting.

10. Battery Compartment – Located on the opposite side of the meter as the LCD screen and Left/Right Function keys, otherwise referred to as the back of the meter. Place that holds batteries.

Methodology of Reagent

When glucose reacts with the reagents on the test strips, an electrical current is produced, which is proportional to the glucose concentration in the blood sample.

The glucose concentration is calculated by the meter type and the amount of electrical current measured.

4. Alternate Site Testing (AST)

What is AST?

Sampling from anatomical sites (parts of the body) other than the fingertip (i.e. forearm, upper arm, thigh, calf, palm) to check the blood glucose levels. This system allows you to test on the forearm with the equivalent results to fingertip testing.

There are important limitations to AST. Please consult your healthcare professional before you use AST.

What is the advantage?

Patients typically feel more pain when lancing their finger pads because they have substantially more nerve endings than other testing sights. It is reported that other testing sites cause less pain; however, there are other limitations. Please speak to your healthcare provider before using this method.

Use Alternate Site Testing (AST) ONLY:

1.	if it has been greater than 2 hours after taking insulin

2.	if it has been greater than 2 hours after a meal

3.	if it has been greater than 2 hours after exercise

*Do not use AST if you are pregnant or if your blood sugars are not stable.

*Do not use AST if you think your blood glucose is low or if your AST results do not match the way you feel.

Testing Procedure for AST

1. Please unscrew the lancing device by turning the cap counter clockwise. A clear cap will come with the package of a meter kit or a lancing device.

1a. Insert the clear cap that comes with your kit.

2. Insert a new lancet firmly into the lancet holder.

3. Twist off the protective top of the lancet.

4. Close the cap of the lancing device and grasp as shown. Slide into locking position. If necessary, please set the lancing device for a deeper puncture.

5. Select a soft, fleshy, area on your forearm that is clean and dry. Make sure it is away from bone and free of visible veins and hair.

6. Massage the selected area gently to increase blood flow to the puncture site.  Clean the test location with soap and water.

7. Open the test strip bottle and take one test strip out.

8. Insert the test strip while the meter is in the OFF position. The meter will turn on with a blinking blood-drop icon on the LCD screen. You will hear: “Hello, Solo is Ready. Apply blood now.”

9. Press and hold the loaded lancing device against the forearm for a few seconds, and press the button.

10. Wait a few seconds until a blood drop forms. Make sure you have sufficient blood to fill the small window of the test strip.

11. Apply your blood to the front edge (the tip) of the test strip. The meter will start counting down from 6 seconds. Once complete you will hear: “Testing in process. The glucose level is number/unit.”

12. Once you remove the used test strip, the meter will turn off automatically. You will hear: “Goodbye.”, and the LCD screen will display “OFF”.

5. Safety Information

Please use this device only for the intended uses described in this user guide.

Before using this system to test your blood glucose, please read instructions thoroughly. Do a quality check on the system following the instructions and consult with your healthcare professionals for questions or problems.

Be aware of the safety of young children or handicapped persons near you when you conduct a glucose test using this system.

Solo-V2TM test strips can only be used with the Solo-V2TM meter.

Please keep the test strip bottle away from children. The test strips and bottle cap can present a choking hazard.

Please be cautious when removing the lancet. Always place the protective cap back on the exposed lancing device.

Never try to disassemble the meter in any circumstance. If your meter is not working properly, or if you need the Technical Support, please contact your local distributor for help.

6. Quick Testing Instructions

Please read through all the detailed instructions before real testing to assure the effectiveness of testing.

1. Wash hands with warm soapy water and dry thoroughly.

2. Remove test strip from bottle and recap the bottle immediately.

3. In the “OFF” position, insert a test strip into the test strip slot.

The meter will automatically turn on, and you will hear “Hello, Solo is Ready. Apply blood now”.

4. Lance the appropriate finger and let a blood drop form.

4a. If you wish to flag the glucose test as a “Before-Meal” test:

·	Keep meter in the “OFF” position

·	Insert NEW strip

·	Listen for “Apply blood now”

·	Press the “Right Function” button two (2) times

·	You are in the right mode when you hear, “ Ready for test Before-Meal”

4b. If you wish to flag the glucose test as an “After-Meal” test:

·	Keep meter in the “OFF” position

·	Insert NEW strip

·	Listen for “Apply blood now”

·	Press the “Right Function” button three (3) times

·	You are in the right mode when you hear, “ Ready for test After-Meal”

4c. If you do not want to flag the test, simply insert the strip when it is in the “OFF” position. After you hear “Apply blood now”, apply blood to the strip. Note: you will have 5 minutes to apply the blood before the meter will automatically turn off.

5. After a 6 second countdown, the test result will appear on the screen. You

will hear: “ The glucose level is number/unit”. For those glucose tests that are flagged, the screen will display an upper-case “A” on the right side of the reading as an After-Meal test and a lower-case “b” as a Before-Meal test.

6. Remove the test strip and the meter will turn off automatically. You will hear“Goodbye” and the LCD screen will display “OFF”.

Notes:

Voice volume can be adjusted by the two Up/Down toggle buttons on the upper right side of the meter.

To change languages of the meter, use the round button on the lower right-hand side of the meter.

7.  Installing the Battery

1. If the meter is on, turn the meter off by pressing and holding the

“Left Function ” button for two seconds.

2. Open the battery cover on the back of the meter by pulling the tab up and away from the meter.

3. Remove used batteries if there are any and replace with two (2) AAA (triple A) alkaline batteries. Put the battery cover back in place and snap it closed. You will hear a “beep” sound if the battery is installed correctly.

Important:

1. It is suggested you replace the batteries when the battery icon appears empty.  You will get an Error 3(E_3) message when the batteries do not have enough power to perform a test.

2. Replacing the batteries will not delete the stored testing results or any other meter setting EXCEPT date and time. You will need to reset the date and time after the batteries have been replaced.

3. Please dispose of used batteries according to your local ordinances.

8. Setting the Audio Voice Volume and Language

On the right side of the meter, you will find two buttons, the “Up” button and the “Down” button. The “Up” button increases volume and the “Down” button decreases volume.

When the meter is in the “OFF” position, the voice volume of the meter can be adjusted by pressing the “Up” or “Down” buttons. There are 8 different volume settings ranging from OFF to a volume level of seven (7). When initially pushing the volume buttons you must press and hold for a second before it will change the volume. This feature prevents the meter from speaking when a button is accidentally touched.

To mute the volume, simply set the meter to setting “ 0 ”, and you will hear “Voice Off ”. The speaker icon will disappear from the LCD screen.

On the right side of the meter you will find a small, round button.

To toggle between English and Spanish simply press the round button twice. The first time you push the button, the meter will speak the current setting. If you push the button a second time, you will switch to the alternate language. The screen will show “E” or “S” and you will hear “English” or “Español”.

If you stop pressing this round button, the meter will automatically turn off in 2 seconds. An audible voice will say, “ Goodbye ” and the LCD screen will display “OFF”.

9. Setting (Date/Time/Alarm/Measuring Unit and Memory Deletion) and

LCD Screen Introduction

Press the “Left Function” button for 2 seconds to enter the “settings” mode.  Once in the “settings” mode the “Right Function” key will be used to toggle between the different settings.  The “Up/Down” on the right side of the meter will be used to increase or decrease the value of current setting.

Below is the appearance of the LCD screen:

(1) Measurement Units

(2) Memory

(3) Setting

(4) A/b/C

(5) Blood drop icon

(6) Battery Symbol

(7) Voice Symbol

(8) Day Average

(9) Alarm Reminder

(10) Month

(11) Day

(12) Hour

(13) Minute

1. Appears with the test result either in mg/dL or in mmol/L.

2. Appears when you recall the memory.

3. Appears when you are in setting mode.

4. Appears as test flags. C for control solution test, b for blood test before meal, A for blood test after meal, no flag for blood test without flag.

5. Flashes when it is ready to apply the sample.

6. Indicates the battery status.

7. Indicates if audio function is on or off.

8. Indicates current displayed result is an average.

9. Appears when alarm is on.

10. Month

11. Day

12. Hour

13. Minute

The LCD screen for Power Off - When the meter is off, the date, time, measurement unit, battery icon, and “OFF” will be still on the screen. If the voice is on or the alarm reminder is on, the speaker icon and the alarm icon will be displayed as well.

To SET the meter:

1. With the meter in the “OFF” position, press the “Left Function” key and hold for 2 seconds. The first item to appear on the LCD screen is a flashing “09”. This represents the last two digits of the current year. The meter is now in the “Setting” mode. You will hear: “You are Now in Setting Mode. Please set the year.”

2. Set the Year - Press the “Up” or“Down” key on the right side of the meter until the last two digits of the desired year are displayed. Each time you adjust the year, you will hear “The year is xxx.” Press the “Right Function” key to confirm the correct year and the meter will move to the next setting.  You will hear “Please set the month.”

3. Set the Month- Press the “Up” or“Down” key on the right side of the meter until the desired month appears. Each time you adjust the month, you will hear “The month is xxx.” Press “Right Function” key to confirm the correct month and the meter will move to the next setting. You will hear:  “Please set the day.”

4. Set the Day - Press the “Up” or“Down” key on the right side of the meter until the desired day appears. Each time you adjust the day, you will hear “The day is the xxx”.  Press “Right Function” key to confirm the correct day and the meter will move to the next setting. You will hear: “Please set the hour.”

5.Set the Hour - Press the “Up” or“Down” key on the right side of the meter until the desired hour appears. Each time you adjust the hour, you will hear: “The hour is xx A.M. or P.M.”  Press “Right Function” key to confirm the correct hour and the meter will move to the next setting. You will hear: “Please set the minute.”

6. Set the Minute - Press the “Up” or “Down” key until the desired minute appears. Each time you adjust the minute, you will hear “The minute is xxx.” Press “Right Function” key to confirm the correct minute and the meter will move to the next setting. You will hear: “Please set the alarm.”

7. Set the Alarm Reminder - Press the “Up” or “Down” key to select the setting of alarm reminder. You will see the alarm icon is flashing with an “AL” display on the LCD screen. You will hear: “Set alarm reminder.” Press “Right Function” key to confirm “Setting mode.”

The hour segment of“00:00” will be flashing on the screen. Press the “Up” or “Down” key until the desired hour appears.  Press the “Right Function ” key to confirm the correct hour for the alarm setting. Each time you adjust the hour, you will hear “The hour is XX A.M. or P.M” if the voice is on. Then the minute segment of“00:00” will be flashing on the screen. Press the “Up” or “Down” key until the desired reminder minute appears on the screen. Each time you adjust the minute, you will hear “The minute is xx.” Press “Right Function” key to confirm the setting and the meter will move to the next setting.

Note: Once the Alarm Reminder is SET, the meter will beep for 60 seconds every day at the set Alarm Reminder time. To silence the alarm, press the “Right Function” key once.  If you have set the alarm to 8:45 am, you will hear: “Alarm Reminder is set to 8:45 A.M.”

If you decide NOT to set the alarm press the “Up” or “Down” key after you first hear: “Please set the alarm.” Press the “Right Function” key to exit the alarm setting. You will hear: “Please set the unit.”

8. Set the measurement unit - Press“Up” or “Down” key on the right side of the meter until the desired measurement appears. The meter will display either “mg/dL” or “mmol/L” when pressing “Up” or “Down” key. Each time you press “Up” or “Down”, you will hear: “Measurement unit in milligrams per deciliter” or“Measurement unit in millimole per liter.” Press “Right Function” key to confirm the selection and you will hear: “Ready to delete memory records. Press the “Down” button to delete all records.

Warning: This meter can show results either in “mg/dL” or “mmol/L” units. Mg/dL is the unit used in the U. S., please contact the manufacturer if your meter does not display mg/dL when you turn it on. The factory set default for the meter is mg/dL, in the event that the meter loses power or the batteries are changed, the factory set default is mg/dL. It is very important to use the correct unit of measurement to properly manage your diabetes. If you live in the United States, your results will never have a decimal point. If you live outside the United States and use the metric system, you should use mmol/L. Your results will always have a decimal point.

9. Delete Memory- You will hear: “Ready to delete memory records. Press the Down button to delete all records”. Press the “Down” button to permanently delete all memory records. You will receive a confirmation “Are you sure you want to delete all records?” Press the “Down” button a second time to delete all records.

If you want to skip this memory deletion step, press the “Right Function” key when you first hear “Ready to delete memory records.” The meter will automatically power off with an audible voice “Goodbye” and the LCD screen will display “OFF”.

10. Running a Control Solution Test

IMPORTANT:

1. Press “Right Function” key for place the meter in “Control Solution” mode. If you leave the meter in the standard mode, the control solution test result will be stored in the memory and affect your average “AVG” results.

1. Insert a strip into the meter, and you will hear: “Hello, Solo is ready.” You will hear: “Apply blood now.”

2. Press the “Right Function” key, and you will see a letter “C” on the in the right corner of the LCD screen to indicate you are in control test mode. You will hear “Apply control solution now.” The test result will not be stored in the memory.

3.  Squeeze a small amount of control solution on a flat surface and apply to the front edge of the test strip. You will hear: “Testing in process.”

4.  Wait 6 seconds and the meter will display the result.  You will hear: “The glucose level is number/unit.”

5.  Remove the used strip. The meter will power off with a display of “OFF” and you will hear: “Goodbye”.

Note:

1. You must press the “Right Function” key only once to set the meter in the control solution mode before running the test. Do not perform the blood test in control solution mode if you want to save the test result in memory.

2. Use only Solo-V2TM control solution with the Solo-V2TM glucose meter only.

3. Solo-V2TM control solutions are used to confirm the accuracy of the meter and the test strips.

4. Solo-V2TM control solution contains a known amount of glucose that reacts with test strips. The Low and High level control solutions are intended to check the monitoring system in different measurement ranges.

5. To ensure an accurate result, make sure you gently roll the control solution in the palm of your hand. Discard the first drop.

6. Use for only 3 months after first opening. Record the open date on the control solution bottle. Discard after 3 months.

7. Compare your control solution test results with the expected range printed on the test strip bottle label.  If your glucose control results fall outside the expected range, repeat the test. Results that repeatedly fall outside the expected range may indicate:

(1) The control solution test may not be performed correctly.

(2) The control solution is expired or contaminated.

(3) The test strips are damaged or expired.

(4) The meter is malfunctioning.

11. Preparing the Lancing Device

1. Unscrew the lancing device by turning the end counter clockwise.

2. Insert a new lancet firmly into the lancet holder.

3. Twist off the protective tip of the lancet.

4. Place the cap back on the end of the lancing device by turning clockwise.

5.. Slide lancing device into locking position.

Important: If you do not obtain enough blood on the strip, you will receive an error message. You will also hear, “Insufficient blood, Please test again.” This is critical for your safety.

11a. Preparing to Test Blood Sugar

1. Wash your hands with warm, soapy, water.

2. Open the test strip bottle and take out one new test strip.

3. Recap the bottle quickly and firmly.

4. Insert a test strip into the meter while the meter is in the “OFF” position. You will hear: “Hello, Solo is ready”.

5. Retrieve the loaded lancing device or single-use lancet.

6. Apply blood to the front edge of the test strip after you hear: “Apply blood now.” You will also see a blinking icon of a blood drop. If enough blood is obtained you will hear: “Testing in process.”

7. To flag a Before-Meal test, press the “Right Function ” twice after you hear “Apply blood now.” Apply blood to the edge of the strip. You will hear “Testing in process”.

8. To flag an After-Meal, press the “Right Function ” three (3) times after you hear “Apply blood now.” Apply blood to the edge of the strip. You will hear “Testing in process”.

9. After a 6 second countdown you will hear: “The glucose level is number/unit”. For those glucose tests that were flagged, the screen will display an “A” as an After-Meal test and a “b” as a Before-Meal test.

10. Once you remove the used strip, the meter will turn off automatically and you will hear: “Goodbye”. The LCD screen will display “OFF”.

12. Recalling the Memory and Viewing Different Averages

Note: The meter will audibly read the latest 10 test results with DATE. You will hear: “September 4th (Month/Day), 110 mg/dL (Glucose value).”

A. Using the Meter Memory

The test results will be stored automatically. The meter can store up to 500 test results at any single point in time. You can also view 7, 14, 28, 60, and 90 averages.

B. Recalling the Test Results from the Meter Memory

1. In the “OFF” position, press “Left Function” key once and you will hear: “Recalling the memory.” If no results are stored, “three dashes” will be displayed on the screen and you will hear: “Zero memory.”

2. Press the “Down” key to retrieve the latest test first.

3. Press the “Up” key to retrieve the oldest test first.

C. Viewing the Averages (AVG) Stored in Memory.

Note: The average will have all test results including those flagged as Before-Meals, After-Meals and those not flagged at all. When the meter displays “---” on the screen, it indicates there is no stored data.

1. Press “Right Function” key once and you will see “AVG” and the number “7” flashing on the lower left corner. The number “7” represents the average of the test results over the last seven days. For example: You will hear: “The average for seven days is eighty-three milligrams per deciliter.”

2. Press “Right Function” key incrementally move between the different averages (7,14, 28, 60, and 90 day averages)

D. Displaying Averages in Sequence

1. Press the “Up” key to see the test averages in the sequence starting with Before-Meal test results.

2. Press the “Down” key to see the test averages in the sequence starting with After-Meal test results.

The screen will display the average reading with a “b” or “A”.

Note:

1. The out-of range HI/LO test results will not be stored in the memory.

2. Do not insert the strip into the meter when you want to recall the test results in memory.

3. The average readings in the meter memory are calculated from the results obtained during the 7, 14, 28, 60 and 90 calendar days preceding the current date and time settings.

13. Expected Values Reference2

Blood glucose levels are fluid and change all the time depending on food intake, medication dosages, health, stress, and exercise. Consult your physician or healthcare professional for your expected glucose levels. Expected plasma blood glucose values for normal, non-diabetic, adults are as follows:

Expected Plasma Blood Glucose for Non-Diabetics

Before Eating(no food intake for 8 hours)    <110mg/dL

Two hours after meals                    <140mg/dL

Consult your physician or healthcare professional for your appropriate range.

If you are experiencing symptoms that are not consistent with your blood glucose results and you have followed all the instructions in this user guide, call your healthcare professional. Never ignore symptoms or make significant changes to your diabetes control programs without speaking to your healthcare professional.

Solo-V2TM meter gives plasma equivalent results.

14. Transmitting Results

Note: Solo-V2TM Glucose Meter allows you to transfer the test results stored in its memory to your personal computer. However, you will need to order software and a Data Download Cable separately from your distributor. The software and its data download procedure will come as a CD in the download cable box. You can also check with your distributor for the software download option. The meter still keeps the results in the memory after transmitting.

The data transmitting procedure is as follows:

1.	Install software from the CD or Download the application from www.SoloMeters.com onto your PC before running the data transmission.

2. Keep meter in the “OFF” position

3. Connect the PC and the meter with the mini USB cable.

4. Hold the “Left Function” and “Right Function” keys down at the same time for 5 seconds until “PC”is displayed and you hear: “Ready to transmit data”.

3.	Press the “Right Function” key to begin transmitting data from your meter to your personal computer.

4.	Press and hold the “Left Function” key for 2 seconds to exit the transmission mode.

15. Maintenance of your System

1. Store your test strips in the bottles provided.

1a. Keep in a clean, dry, place between 46-86 o F (8~30oC)

2. Clean the meter after every use. Gently wipe and clean surface of the meter with a soft cloth that has been slightly dampened with alcohol.

3. Handle with care and do not drop the meter.

4. Enforce infection control policies when being used by healthcare professionals.

We suggest you should periodically compare the test system to another test system which is well maintained and monitored by a healthcare provider.

16. Troubleshooting

Following is a summary of all Error Messages. These messages help to identify certain problems, but do not appear in all cases when a problem has occurred. Improper use may cause an inaccurate result without producing an error message or a symbol. In the event of a problem, refer to the information in the table under Solution.

Error messages:

E_1: The meter temperature is below its operating range.  (Audible message: “The temperature is too low.”)

E_2: The meter temperature is above its operating range.  (Audible message: “The temperature is too high.”)

E_3: Battery Low. No further tests can be performed. (Audible message: “Battery low”)

E_4: The blood sample is not sufficient. (Audible message: “No sufficient blood ”)

E_5: The strip is wet or used. (Audible message: “Strip is wet or used.”)

E_6: The meter is not working properly, or the test strip is damaged. (Audible message: “Error in meter or strip”)

E_7: Memory damaged. (Audible message: “Memory damaged.”)

Other error messages:

H I: Test result is higher than 600 mg/dL. (Audible message: “The glucose level is too high”)

LO: Test result is lower than 20mg/dL. (Audible message: “The glucose level is too low”)

Error Messages

| Message | Audible Message | Possible Cause | Solution |
| --- | --- | --- | --- |
| E_1 | The temperature is too low. | The meter is operating in an ambient temperature below 10℃ or 50℉ which is the lowest range of operating temperature. | System operation range is 10℃~40℃ (50℉~ 104℉). Repeat the test when the temperature is within the range. |
| E_2 | The temperature is too high. | The meter is operating in an ambient temperature above 40℃ or 104℉ which is the cap of operating temperature. |  |
| E_3 | Battery low | The battery is too low to operate the meter. | Replace the battery immediately. |
| E_4 | No sufficient blood | The blood sample is not sufficient. | Please re-test with a new strip and a new blood sample. |
| E_5 | Strip is wet or used. | The inserted test strip has been wet or used. | Re-test with a new test strip and a new blood sample. |
| E_6 | Error in meter or strip | The meter is not working properly because of a defective meter or a defective strip. | Re-test with a new strip. If error message E_6 is still displayed, contact your local distributor for technical support. |
| E_7 | Memory damaged | The memory chip of the meter could be damaged or malfunctioning. | Re-test with a new strip. If error message E_7 is still displayed, contact your local distributor for technical support. |

Special Messages

| Message | Audible Message | Possible Cause | Solution |
| --- | --- | --- | --- |
| LO | Test result is lower than 20 mg/dL (1.1 mmol/L) | The test result is lower than the measure range of 20 mg/dL (1.1 mmol/L). | Test again following the user guide for correct glucose measurement process. If you see LO again, please call your medical doctor for advice immediately. |
| HI | Test result is higher than 600 mg/dL (33.3mmol/L) | The test result is lower than the measure range of 600 mg/dL (33.3 mmol/L). | Test again following the user guide for correct glucose measurement process. If you see HI again, please call your medical doctor for advice immediately. |

Problem in Operation

| Problem | Possible Cause | Solution |
| --- | --- | --- |
| The meter will not turn on when a strip is inserted | 1) The meter is defective or the battery is low 2) The strip is not inserted properly or the strip is damaged. | 1) Make sure the arrow on the inserted strip is facing towards the meter, and the blood icon on the strip is facing up. 2) Take a new strip and re-insert into the meter. 3) Check the battery power or replace them if necessary. 4) Call Customer Service. |
| The meter does not count down | The meter is defective or the battery is low | Replace the battery. If the meter fails to operate, call Technical Support. |
| The meter display turns off abnormally | The meter is defective or the battery is low | Replace the battery. If the meter fails to operate, call Technical Support. |
| The control test result is out of specified range. | The control solution or test strip maybe outdated or defective. | Verify if the control solution or the test strip is expired. If they are not expired, please call Customer Service |

Warning: Your meter can show results either in “mg/dL” or “mmol/L” units. Mg/dL is the unit used in the U. S., please contact the manufacturer if your meter does not display mg/dL when you turn it on. The factory set default for the meter is mg/dL, in the event that the meter loses power or the batteries are changed, the factory set default is mg/dL. It is very important to use the correct unit of measurement to properly manage your diabetes. If you live in the U.S.A. you should use mg/dL. Your results will never have a decimal point. If you live outside the United States and use the metric system, you should use mmol/L. Your result will always have a decimal point.

17. Customer Service

If you need assistance with your Solo-V2TM Glucose Monitoring System, please contact:

BIOSENSE MEDICAL DEVICES, LLC

6555 Sugarloaf Parkway,

Suite 307-168,

Duluth, GA 30097

TEL: 877-592-3922

FAX: 888-620-2345

(9 a.m.~5 p.m. US EST/ Mon.~Fri. Outside these times, contact your healthcare professional)

For questions related to your health condition, please call your doctor or healthcare professional.

18. Limitations

1. Do not use for neonatal blood glucose testing.

2. Hematocrits in the range of 32~56% do not affect the blood glucose results. If you do not know your hematocrit level, consult your healthcare professional.

3. Hemoglobin levels of 500 mg/dL or below will not interfere with blood glucose test results.

4. Cholesterol levels up to 500 mg/dL (12.9 mmol/L) and Triglycerides up to 1000 mg/dL has been shown not to affect glucose results.

5. Medications

Interference was observed for therapeutic levels of L-Dopa. No interference was shown for uric acid, acetaminophen, ascorbic acid and ibuprofen in normal therapeutic levels.. However, higher concentrations in blood may cause incorrect results:

Uric acid: ＞10.9 mg/dL; acetaminophen: ＞6.2 mg/dL;

ascorbic acid: ＞4.5 mg/dL; ibuprofen: ＞37.5 mg/dL.

6. The test strips may be used at altitudes up to 5,280 feet (1,609m) without an effect on test results.

7. Persons suffering from severer dehydration should not be tested using a capillary whole blood sample.

8. Test results below 70 mg/dL indicate low blood glucose (hypoglycemia). Test results greater than 240 mg/dL indicate high blood glucose (hyperglycemia). If you get results below 70 mg/dL or above 240 mg/dL, repeat the test, and if the results are still below 70 mg/dL or above 240 mg/dL, please consult your healthcare professional immediately.2

9. Inaccurate results may occur in severely hypotensive individuals or patients in shock. Inaccurate results may occur for individuals experiencing a hyperglycemic-hyperosmolar state, with or without ketosis. Critically ill patients should not be tested with a blood glucose meter.

19. Specifications

Test: Glucose

Sample: Whole Blood

Principle of the test method: Amperometric, glucose oxidase.

Test sites : Fingertip, Forearm

Measuring Range: 20~600 mg/dL (1.1~33.3 mmol/L).

Measuring Time: 6 seconds

Code required: No code device

Batteries: ALKALINE 24A  LR03  AAA 1.5V X 2)

Operating Range: 10~40℃(50-104℉)

Humidity: 20~80 % RH

Width: 55 mm.

Length: 100 mm.

Thickness: 18 mm.

Weight:~ 75g.

Memory: 500 test results with date and time.

Power: Automatic shut-off after 5 minutes when a strip is inserted, or 20 seconds if no strip is inserted.

NDC #: 8611-501001

The device has certified to meet the following standard:

98/79/EC, EN 60601-1 + EN 60601-1-1

ISO 15197

Note: Please refer to the package insert that came with your test strips. You will find the information on how the system works

20. Ordering Supplies

To order supplies, please contact:

BIOSENSE MEDICAL DEVICES, LLC

6555 Sugarloaf Parkway,

Suite 307-168,

Duluth, GA 30097

TEL: 877-592-3922

FAX: 888-620-2345

(9 a.m.~5 p.m. US EST/ Mon.~Fri. Outside these times, contact your healthcare professional)

21. Warranty Information

The Solo-V2TM Glucose Meter is warranted to be free from defects for one year from date of purchase. For warranty assistance, please contact:

BIOSENSE MEDICAL DEVICES, LLC

6555 Sugarloaf Parkway,

Suite 307-168,

Duluth, GA 30097

TEL: 877-592-3922

FAX: 888-620-2345

(9 a.m.~5 p.m. US EST/ Mon.~Fri. Outside these times, contact your healthcare professional)

Reference:

1. American Diabetes Association：Diabetes Care, January 2007, volume 30 (Suppl. 1) S42-S47

2. American Diabetes Association-Diabetes Forecast (website information http://www.forecast. Diabetes.org/diabetes-101/hyperglycemia. http://www.forecast.diabetes.org/diabetes-101/hypoglycemia.)